Servo Controller command-response protocol
All command messages must be terminated with a two-byte CRC code followed by ETX {0x03}. All controller response messages are terminated with a two-byte CRC code followed by ETX. Except as a terminating character, and possibly as a byte in the CRC code, ETX may not appear in any message.

Command format details
Messages may contain multiple commands, separated by whitespace (spaces, tabs, new lines, or carriage returns).
[n] indicates that a whole number is expected.
[d] indicates that a decimal number is expected. A whole number is acceptable. Any digits provided beyond the expected precision are ignored. For example, "3", "-4.7", and "134.00" are all acceptable decimal inputs.

Response format details
In controller report templates, '#' represents a numeric digit, '.', a decimal point, '-', a negative symbol. Other template symbols represent characters that may be non-numeric. When a reported number is positive, the negative symbol is replaced with a space, so the message length remains fixed.

Command	Description
r	Request a status report.

Response format template
"### ##### # ## ## #### #### ###.## ###.## ##.### #####" (54 bytes + "\r\n")

The values represented by the symbols in the template above are, in sequence:
Servo (the currently selected servo channel)
ServoCpw (servo control pulse width)
CPEnabled (control pulses are enabled, i.e., whether servo motor is on (1) or off (0))
StopOnLim0, Limit0 (whether limit0 stops the servo (0 or 1), and whether it is engaged (0 or 1))
StopOnLim1, Limit1 (whether limit1 stops the servo (0 or 1), and whether it is engaged (0 or 1))	StopOnI (the current limit in milliamps; if exceeded, the servo stops. 0 means no current limit)
I (servo current in milliamps)
StopOnT (time limit; servo runs this many seconds, max. 0 means no time limit)
Elapsed (seconds the servo has been running, or was running before it last stopped)
V (servo power supply voltage)
Error (sum of active codes from the following list)
0	No error
1	ADC out of range
2	RS232 input buffer overflow
4	RS232 CRC error
8	Unrecognized command from RS232
16	Invalid device channel
32	Datalogging time interval out of range (0..255)
64	Control pulse width out of range
128	StopOnI out of range
256	StopOnT out of range
512	Unrecognized stop limit
1024	Both limit switches activated?
2048	Low servo power supply voltage

r[n]	Start datalogging, or stop, if n is 0. Issue a controller status report every n seconds. [0..255]
n[n]	Select the servo connected to channel n. [0..63]

p[n]	Set the servo's control pulse width (ServoCpw) to n microseconds [CPW_MIN..CPW_MAX]

g	Go. Move the servo to the commanded position (ServoCpw). (CPEnabled = 1)

g[n]	Set ServoCpw and "go". [CPW_MIN..CPW_MAX]

0 (zero)	Position the servo to the center of its movement range (ServoCpw = 1500 microseconds).

s	Stop any servo movement.

c	Clears the I (servo current) and Elapsed time values (sets them to 0).

h	Show a report header.

Response
"SRV __CPW G L0 L1 ILIM ___I __TLIM __ELAP _____V Error\r\n" (54 bytes + "\r\n")

i[n]	Set StopOnI (current limit in milliamps) [0..32767], but the max practical value is < 1600
If n is 0, current limit is disabled.

t[d]	Set StopOnT (time limit in seconds) [0.00..300.00]
If n is 0, time limit is disabled.

l[d]	(lower case 'L', not the number '1') Set Stop Limit [10, -10, 11, -11]
10 == enable limit 0
-10 == disable limit 0
11 == enable limit 1
-11 == disable limit 1

z	Show program data.

Response format template
"Aeon Laboratories SC64 V.########-####" (38 bytes + "\r\n")
"S/N:####" (8 bytes + "\r\n")
"CPW_MIN:#### CPW_MAX:######" (29 bytes + "\r\n")

The first line contains the servo controller model number and firmware version. The second line has the serial number. The third line shows the limits for the servo control pulse width ServoCpw, in microseconds.